## Cover
From the Laboratories at
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel 603-715-3000 • Fax: 603-225-3089
CALIBRATION CERTIFICATE
This calibration is traceable to the National Institute of Standards and Technology in accordance with ISO 10012-1 and MIL-STD-45662. The calibration was accomplished by comparison to standards maintained by the laboratories at SphereOptics LLC. Complete compliance records and procedures are maintained by SphereOptics LLC and are available for inspection upon request.
CUSTOMER: | University of Maryland
1000 Hilltop Circle
BALTIMORE, MD 21250
USA
PURCHASE ORDER #: | UMBC1-0000005227
MODEL #: | SR-3BC
SERIAL #: | 5585
ITEM DESCRIPTION: | SPECTRAL RADIANCE CALIBRATION 300 - 2400nm
As Received |  | As Left
| New Equipment |  | In Tolerance*
| Out of Tolerance* |  | Out of Tolerance*
| Inoperative |  | Inoperative
| Limited Calibration** |  | Calibration
| Used Equipment |  | Adjusted
* Per Referenced calibration procedure
** Limited by customer of equipment performance
This report or any attached document within this report shall not be reproduced, except in full, without the written permission of SphereOptics LLC.
CALIBRATED: | 2008-07-29 | TECHNICIAN: | Jjablonski
AUTHORIZED SIGNATURE:
RECALIBRATION DUE: | 2009-07-29
VP Engineering

## Data
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Fax: 603-225-3089
www.sphereoptics.com
CUSTOMER: |  | University of Maryland
MODEL NUMBER: |  | SR-3BC
SERIAL NUMBER: |  | 5585
DATE: |  | 2008-07-29
SPECTRAL RADIANCE DATA
EQUIPMENT USED
|  |  | Lamp | Address |  | Current (A) | Volts (V)
Agilent | 3634A |  | ETHS-150 | 5 |  | 5.79 | 19.52
Agilent | 5566 |  | ITHS-300 | 6 |  | 0 | 0
Agilent | 3634A |  | ITHS-150 | 7 |  | 0 | 0
Agilent | 3634A |  | ITHS-75 | 8 |  | 0 | 0
Kiethley | 6485 | [5 empty] | 1.225e-05 | A
SOH-1 | Radiance Standard
* Above PS Address must be the same for LR software to work correctly
Wavelength | Radiance |  |  | Wavelength | Radiance |  | Wavelength | Radiance
nm |  | uW/cm2-sr-nm |  | nm | uW/cm2-sr-nm |  | nm | uW/cm2-sr-nm
300 |  | 0.066178 |  | 580 | 5.8145 |  | 860 | 9.7571
310 |  | 0.098836 |  | 590 | 6.0897 |  | 870 | 9.8321
320 |  | 0.13748 |  | 600 | 6.3824 |  | 880 | 9.8896
330 |  | 0.22579 |  | 610 | 6.639 |  | 890 | 9.9961
340 |  | 0.23718 |  | 620 | 6.9092 |  | 900 | 10.082
350 |  | 0.31774 |  | 630 | 7.1592 |  | 910 | 10.104
360 |  | 0.41963 |  | 640 | 7.4022 |  | 920 | 10.1681
370 |  | 0.52354 |  | 650 | 7.6263 |  | 930 | 10.2096
380 |  | 0.6425 |  | 660 | 7.8383 |  | 940 | 10.1989
390 |  | 0.79733 |  | 670 | 8.0534 |  | 950 | 10.2276
400 |  | 0.95148 |  | 680 | 8.224 |  | 960 | 10.2092
410 |  | 1.1826 |  | 690 | 8.3962 |  | 970 | 10.2349
420 |  | 1.3536 |  | 700 | 8.569 |  | 980 | 10.2044
430 |  | 1.527 |  | 710 | 8.7251 |  | 990 | 10.198
440 |  | 1.7786 |  | 720 | 8.8548 |  | 1000 | 10.2172
450 |  | 2.0092 |  | 730 | 8.9879 |  | 1010 | 10.1709
460 |  | 2.2755 |  | 740 | 9.079 |  | 1020 | 10.1536
470 |  | 2.5258 |  | 750 | 9.1989 |  | 1030 | 10.1228
480 |  | 2.7992 |  | 760 | 9.2369 |  | 1040 | 10.0992
490 |  | 3.0963 |  | 770 | 9.2831 |  | 1050 | 10.0312
500 |  | 3.3776 |  | 780 | 9.3782 |  | 1060 | 9.9429
510 |  | 3.6937 |  | 790 | 9.4068 |  | 1070 | 9.8695
520 |  | 3.997 |  | 800 | 9.4573 |  | 1080 | 9.7736
530 |  | 4.2998 |  | 810 | 9.4925 |  | 1090 | 9.6925
540 |  | 4.5961 |  | 820 | 9.5279 |  | 1100 | 9.5727
550 |  | 4.8952 |  | 830 | 9.5855 |  | 1110 | 9.4409
560 |  | 5.2117 |  | 840 | 9.6264 |  | 1120 | 9.2514
570 |  | 5.5231 |  | 850 | 9.6793 |  | 1130 | 9.1685
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Tel: 603-225-3089
Wavelength | Radiance |  |  | Wavelength | Radiance |  | Wavelength | Radiance
nm |  | uW/cm2-sr-nm |  | nm | uW/cm2-sr-nm |  | nm | uW/cm2-sr-nm
1140 |  | 9.0214 |  | 1590 | 3.794 |  | 2040 | 1.1283
1150 |  | 8.9039 |  | 1600 | 3.7621 |  | 2050 | 1.0908
1160 |  | 8.7657 |  | 1610 | 3.7122 |  | 2060 | 1.1248
1170 |  | 8.6419 |  | 1620 | 3.6645 |  | 2070 | 1.1006
1180 |  | 8.4847 |  | 1630 | 3.6102 |  | 2080 | 1.0718
1190 |  | 8.376 |  | 1640 | 3.5629 |  | 2090 | 1.0668
1200 |  | 8.2393 |  | 1650 | 3.5106 |  | 2100 | 1.0722
1210 |  | 8.1263 |  | 1660 | 3.4578 |  | 2110 | 1.0715
1220 |  | 8.0523 |  | 1670 | 3.4001 |  | 2120 | 1.0472
1230 |  | 7.9603 |  | 1680 | 3.3315 |  | 2130 | 1.0568
1240 |  | 7.865 |  | 1690 | 3.2512 |  | 2140 | 1.0374
1250 |  | 7.7636 |  | 1700 | 3.157 |  | 2150 | 1.0383
1260 |  | 7.6892 |  | 1710 | 3.0368 |  | 2160 | 1.0316
1270 |  | 7.5692 |  | 1720 | 2.9892 |  | 2170 | 1.0114
1280 |  | 7.4561 |  | 1730 | 2.9097 |  | 2180 | 0.94771
1290 |  | 7.2963 |  | 1740 | 2.8702 |  | 2190 | 0.92871
1300 |  | 7.1549 |  | 1750 | 2.7667 |  | 2200 | 0.93641
1310 |  | 7.0619 |  | 1760 | 2.6931 |  | 2210 | 0.90728
1320 |  | 6.9223 |  | 1770 | 2.685 |  | 2220 | 0.92351
1330 |  | 6.7591 |  | 1780 | 2.6344 |  | 2230 | 0.90828
1340 |  | 6.5338 |  | 1790 | 2.5168 |  | 2240 | 0.84912
1350 |  | 6.3362 |  | 1800 | 2.5014 |  | 2250 | 0.81468
1360 |  | 6.1199 |  | 1810 | 2.4141 |  | 2260 | 0.92382
1370 |  | 5.9216 |  | 1820 | 2.3752 |  | 2270 | 0.72169
1380 |  | 5.6442 |  | 1830 | 2.2785 |  | 2280 | 0.73399
1390 |  | 5.2589 |  | 1840 | 2.264 |  | 2290 | 0.85043
1400 |  | 4.8077 |  | 1850 | 2.1607 |  | 2300 | 0.73855
1410 |  | 4.4944 |  | 1860 | 2.0619 |  | 2310 | 0.68077
1420 |  | 4.2768 |  | 1870 | 1.8842 |  | 2320 | 0.61431
1430 |  | 4.1129 |  | 1880 | 1.6649 |  | 2330 | 0.71055
1440 |  | 3.9693 |  | 1890 | 1.3913 |  | 2340 | 0.60316
1450 |  | 3.8933 |  | 1900 | 1.1831 |  | 2350 | 0.5827
1460 |  | 3.8295 |  | 1910 | 1.0618 |  | 2360 | 0.53691
1470 |  | 3.8314 |  | 1920 | 0.94876 |  | 2370 | 0.52123
1480 |  | 3.8752 |  | 1930 | 0.93538 |  | 2380 | 0.49656
1490 |  | 3.892 |  | 1940 | 0.9264 |  | 2390 | 0.49839
1500 |  | 3.9444 |  | 1950 | 0.96862 |  | 2400 | 0.56432
1510 |  | 3.9587 |  | 1960 | 0.96853
1520 |  | 3.986 |  | 1970 | 0.98659
1530 |  | 3.9775 |  | 1980 | 1.0147
1540 |  | 3.9743 |  | 1990 | 1.0776
1550 |  | 3.963 |  | 2000 | 1.1271
1560 |  | 3.9221 |  | 2010 | 1.1392
1570 |  | 3.8897 |  | 2020 | 1.1416
1580 |  | 3.8249 |  | 2030 | 1.1484
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Tel: 603-225-3089
NIST CALIBRATION TRANSFER UNCERTAINTY
Transfer uncertainty of our secondary standards relative to the National Institute of Standards and Technology (NIST) scale is estimated as follows:
Wavelength (nm) |  | NIST Spectral Radiance Standard |  | SOH transfer uncertianty | Total Uncertianty
|  | k=2 (%) |  | k=2(%)
400 |  | 1.09 |  | 0.63 | 1.259
500 |  | 0.85 |  | 0.44 | 0.95713
800 |  | 0.68 |  | 0.33 | 0.75584
900 |  | 0.52 |  | 0.29 | 0.5954
1000 |  | 0.47 |  | 0.32 | 0.56859
1100 |  | 1.18 |  | 0.56 | 1.3061
1500 |  | 1.18 |  | 0.58 | 1.3148
2000 |  | 1.75 |  | 0.71 | 1.8885
2400 |  | 2.6 |  | 8.86 | 9.2336
This accuracy statement assumes similar geometric use of the standards by the customer and does not account for drift with time and use for the individual items issued.

## Raw
nm | radiance(uW/cm2-sr-nm)
300 | 0.066178
310 | 0.098836
320 | 0.13748
330 | 0.22579
340 | 0.23718
350 | 0.31774
360 | 0.41963
370 | 0.52354
380 | 0.6425
390 | 0.79733
400 | 0.95148
410 | 1.1826
420 | 1.3536
430 | 1.527
440 | 1.7786
450 | 2.0092
460 | 2.2755
470 | 2.5258
480 | 2.7992
490 | 3.0963
500 | 3.3776
510 | 3.6937
520 | 3.997
530 | 4.2998
540 | 4.5961
550 | 4.8952
560 | 5.2117
570 | 5.5231
580 | 5.8145
590 | 6.0897
600 | 6.3824
610 | 6.639
620 | 6.9092
630 | 7.1592
640 | 7.4022
650 | 7.6263
660 | 7.8383
670 | 8.0534
680 | 8.224
690 | 8.3962
700 | 8.569
710 | 8.7251
720 | 8.8548
730 | 8.9879
740 | 9.079
750 | 9.1989
760 | 9.2369
770 | 9.2831
780 | 9.3782
790 | 9.4068
800 | 9.4573
810 | 9.4925
820 | 9.5279
830 | 9.5855
840 | 9.6264
850 | 9.6793
860 | 9.7571
870 | 9.8321
880 | 9.8896
890 | 9.9961
900 | 10.082
910 | 10.104
920 | 10.1681
930 | 10.2096
940 | 10.1989
950 | 10.2276
960 | 10.2092
970 | 10.2349
980 | 10.2044
990 | 10.198
1000 | 10.2172
1010 | 10.1709
1020 | 10.1536
1030 | 10.1228
1040 | 10.0992
1050 | 10.0312
1060 | 9.9429
1070 | 9.8695
1080 | 9.7736
1090 | 9.6925
1100 | 9.5727
1110 | 9.4409
1120 | 9.2514
1130 | 9.1685
1140 | 9.0214
1150 | 8.9039
1160 | 8.7657
1170 | 8.6419
1180 | 8.4847
1190 | 8.376
1200 | 8.2393
1210 | 8.1263
1220 | 8.0523
1230 | 7.9603
1240 | 7.865
1250 | 7.7636
1260 | 7.6892
1270 | 7.5692
1280 | 7.4561
1290 | 7.2963
1300 | 7.1549
1310 | 7.0619
1320 | 6.9223
1330 | 6.7591
1340 | 6.5338
1350 | 6.3362
1360 | 6.1199
1370 | 5.9216
1380 | 5.6442
1390 | 5.2589
1400 | 4.8077
1410 | 4.4944
1420 | 4.2768
1430 | 4.1129
1440 | 3.9693
1450 | 3.8933
1460 | 3.8295
1470 | 3.8314
1480 | 3.8752
1490 | 3.892
1500 | 3.9444
1510 | 3.9587
1520 | 3.986
1530 | 3.9775
1540 | 3.9743
1550 | 3.963
1560 | 3.9221
1570 | 3.8897
1580 | 3.8249
1590 | 3.794
1600 | 3.7621
1610 | 3.7122
1620 | 3.6645
1630 | 3.6102
1640 | 3.5629
1650 | 3.5106
1660 | 3.4578
1670 | 3.4001
1680 | 3.3315
1690 | 3.2512
1700 | 3.157
1710 | 3.0368
1720 | 2.9892
1730 | 2.9097
1740 | 2.8702
1750 | 2.7667
1760 | 2.6931
1770 | 2.685
1780 | 2.6344
1790 | 2.5168
1800 | 2.5014
1810 | 2.4141
1820 | 2.3752
1830 | 2.2785
1840 | 2.264
1850 | 2.1607
1860 | 2.0619
1870 | 1.8842
1880 | 1.6649
1890 | 1.3913
1900 | 1.1831
1910 | 1.0618
1920 | 0.94876
1930 | 0.93538
1940 | 0.9264
1950 | 0.96862
1960 | 0.96853
1970 | 0.98659
1980 | 1.0147
1990 | 1.0776
2000 | 1.1271
2010 | 1.1392
2020 | 1.1416
2030 | 1.1484
2040 | 1.1283
2050 | 1.0908
2060 | 1.1248
2070 | 1.1006
2080 | 1.0718
2090 | 1.0668
2100 | 1.0722
2110 | 1.0715
2120 | 1.0472
2130 | 1.0568
2140 | 1.0374
2150 | 1.0383
2160 | 1.0316
2170 | 1.0114
2180 | 0.94771
2190 | 0.92871
2200 | 0.93641
2210 | 0.90728
2220 | 0.92351
2230 | 0.90828
2240 | 0.84912
2250 | 0.81468
2260 | 0.92382
2270 | 0.72169
2280 | 0.73399
2290 | 0.85043
2300 | 0.73855
2310 | 0.68077
2320 | 0.61431
2330 | 0.71055
2340 | 0.60316
2350 | 0.5827
2360 | 0.53691
2370 | 0.52123
2380 | 0.49656
2390 | 0.49839
2400 | 0.56432