## Cover
From the Laboratories at
SphereOptics, LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel 603-715-3000 • Fax: 603-225-3089
CALIBRATION CERTIFICATE
This calibration is traceable to the National Institute of Standards and Technology in accordance with ISO 10012-1 and MIL-STD-45662. The calibration was accomplished by comparison to standards maintained by the laboratories at SphereOptics LLC. Complete compliance records and procedures are maintained by SphereOptics LLC and are available for inspection upon request.
CUSTOMER: | University of Maryland
1000 Hilltop Circle
BALTIMORE, MD 21250
USA
PURCHASE ORDER #: | UMBC1-0000005227
MODEL #: | UM-S and UM-A
SERIAL #: | 5585
ITEM DESCRIPTION: | Spatial and Angular Uniformity Mapping
As Received |  | As Left
| New Equipment |  | In Tolerance*
| Out of Tolerance* |  | Out of Tolerance*
| Inoperative |  | Inoperative
| Limited Calibration** |  | Calibration
| Used Equipment |  | Adjusted
* Per Referenced calibration procedure
** Limited by customer of equipment performance
TEST PROCEDURE: | Spatial and Angular Uniformity Mapping
EQUIPMENT USED:
This report or any attached document within this report shall not be reproduced, except in full, without the written permission of SphereOptics, LLC.
CALIBRATED: | 2008-07-29 | TECHNICIAN: | JJacobs
AUTHORIZED SIGNATURE:
RECALIBRATION DUE: | 2009-07-29
Calibration Lab Manager

## 1. uniformity
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel 603-715-3000
Fax: 603-225-3089
DATE: |  | 2008-07-29
CUSTOMER: |  | University of Maryland
ITEM: |  | UM-S and UM-A
SALES ORDER |  | 5585
PURCHASE ORDER |  | UMBC1-0000005227
SERIAL NUMBER: |  | 5585
SPATIAL and ANGULAR UNIFORMITY
Equipment Used: Minolta LS-110 luminance meter
**Mapping Performed at Full Luminance Level
353
|  | -0.8 | -0.6 | -0.4 | -0.2 | 0 | 0.2 | 0.4 | 0.6 | 0.8
| 0.8 |  |  |  | 352 | 353 | 353
| 0.6 |  |  | 353 | 353 | 353 | 353 | 353
| 0.4 |  | 354 | 353 | 353 | 353 | 353 | 353 | 353
| 0.2 | 354 | 353 | 353 | 353 | 353 | 353 | 353 | 353 | 352
| 0 | 354 | 353 | 353 | 353 | 353 | 353 | 353 | 353 | 352
| -0.2 | 355 | 354 | 353 | 353 | 353 | 353 | 353 | 353 | 352
| -0.4 |  | 353 | 352 | 352 | 352 | 352 | 352 | 352
| -0.6 |  |  | 353 | 352 | 352 | 352 | 351
| -0.8 |  |  |  | 351 | 352 | 351
|  | -3.2 | -2.4 | -1.6 | -0.8 | 0 | 0.8 | 1.6 | 2.4 | 3.2
| 3.2 |  |  | 98.2% | 98.35% | 98.43% | 98.43% | 98.35%
| 2.4 |  | 98.43% | 98.59% | 98.74% | 98.82% | 98.82% | 98.74% | 98.74%
| 1.6 | 98.66% | 98.9% | 98.98% | 99.05% | 99.13% | 99.29% | 99.52% | 99.37% | 99.37%
| 0.8 | 99.21% | 99.29% | 99.45% | 99.45% | 99.52% | 99.84% | 100.1% | 100.1% | 99.99%
| 0 | 99.6% | 99.68% | 99.91% | 99.99% | 100.1% | 100.2% | 100.5% | 100.6% | 100.5%
| -0.8 | 99.84% | 99.99% | 100.3% | 100.4% | 100.5% | 100.6% | 100.9% | 100.9% | 100.9%
| -1.6 | 100.1% | 100.3% | 100.5% | 100.7% | 100.9% | 101% | 101.2% | 101.1% | 101.2%
| -2.4 |  | 100.7% | 100.9% | 101.2% | 101.3% | 101.5% | 101.4% | 101.3%
| -3.2 |  |  | 101.2% | 101.4% | 101.6% | 101.6% | 101.6%
[5 empty] | -30 | 0 | 30
|  |  |  | 30 | 96.71% | 96.94% | 100.1%
|  |  |  | 0 | 98.38% | 98.84% | 101.3%
|  |  |  | -30 | 101.3% | 102.5% | 103.8%
Horizontal position across exit port
Vertical position across exit port
Angular position from optic axis of port
Angular position from optic axis of port

## Data
| -3.2 | -2.4 | -1.6 | -0.8 | 0 | 0.8 | 1.6 | 2.4 | 3.2 | detector
3.2 |  |  | 12570 | 12590 | 12600 | 12600 | 12590 |  |  | 1 |  |  |  | 12570 | 12590 | 12600 | 12600 | 12590
2.4 |  | 12600 | 12620 | 12640 | 12650 | 12650 | 12640 | 12640 |  | 1 |  |  | 12600 | 12620 | 12640 | 12650 | 12650 | 12640 | 12640
1.6 | 12630 | 12660 | 12670 | 12680 | 12690 | 12710 | 12740 | 12720 | 12720 | 1 |  | 12630 | 12660 | 12670 | 12680 | 12690 | 12710 | 12740 | 12720 | 12720
0.8 | 12700 | 12710 | 12730 | 12730 | 12740 | 12780 | 12810 | 12810 | 12800 | 1 |  | 12700 | 12710 | 12730 | 12730 | 12740 | 12780 | 12810 | 12810 | 12800
0 | 12750 | 12760 | 12790 | 12800 | 12810 | 12830 | 12860 | 12880 | 12870 | 1 |  | 12750 | 12760 | 12790 | 12800 | 12810 | 12830 | 12860 | 12880 | 12870
-0.8 | 12780 | 12800 | 12840 | 12850 | 12870 | 12880 | 12910 | 12920 | 12920 | 1 |  | 12780 | 12800 | 12840 | 12850 | 12870 | 12880 | 12910 | 12920 | 12920
-1.6 | 12820 | 12840 | 12870 | 12890 | 12920 | 12930 | 12960 | 12940 | 12950 | 1 |  | 12820 | 12840 | 12870 | 12890 | 12920 | 12930 | 12960 | 12940 | 12950
-2.4 |  | 12890 | 12920 | 12950 | 12970 | 12990 | 12980 | 12970 |  | 1 |  |  | 12890 | 12920 | 12950 | 12970 | 12990 | 12980 | 12970
-3.2 |  |  | 12950 | 12980 | 13000 | 13010 | 13000 |  |  | 1 |  |  |  | 12950 | 12980 | 13000 | 13010 | 13000
12801.0145
| -3.2 | -2.4 | -1.6 | -0.8 | 0 | 0.8 | 1.6 | 2.4 | 3.2
3.2 |  |  | 98.2% | 98.35% | 98.43% | 98.43% | 98.35%
2.4 |  | 98.43% | 98.59% | 98.74% | 98.82% | 98.82% | 98.74% | 98.74%
1.6 | 98.66% | 98.9% | 98.98% | 99.05% | 99.13% | 99.29% | 99.52% | 99.37% | 99.37% | average | 100%
0.8 | 99.21% | 99.29% | 99.45% | 99.45% | 99.52% | 99.84% | 100.1% | 100.1% | 99.99% | span | 3.437%
0 | 99.6% | 99.68% | 99.91% | 99.99% | 100.1% | 100.2% | 100.5% | 100.6% | 100.5% | +/- | 1.719%
-0.8 | 99.84% | 99.99% | 100.3% | 100.4% | 100.5% | 100.6% | 100.9% | 100.9% | 100.9%
-1.6 | 100.1% | 100.3% | 100.5% | 100.7% | 100.9% | 101% | 101.2% | 101.1% | 101.2%
-2.4 |  | 100.7% | 100.9% | 101.2% | 101.3% | 101.5% | 101.4% | 101.3%
-3.2 |  |  | 101.2% | 101.4% | 101.6% | 101.6% | 101.6%
|  |  | -30 | 0 | 30
|  | 30 | 12750 | 12780 | 13200
|  | 0 | 12970 | 13030 | 13360
|  | -30 | 13360 | 13510 | 13690
13183.3333
|  |  | -30 | 0 | 30
|  | 30 | 96.71% | 96.94% | 100.1% | Span | 7.13%
|  | 0 | 98.38% | 98.84% | 101.3% | +/- | 3.565%
|  | -30 | 101.3% | 102.5% | 103.8%