University of Maryland Baltimore County - UMBC

Phys650 - Special Topics in Experimental Atmospheric Physics (Spring 2009)

J. Vanderlei Martins and Manfredo H. Tabacniks

http://userpages.umbc.edu/~martins/PHYS650/

Instructions for Sphere Calibration Files:

Make sure you wear UV protective glasses whenever you are working with the sphere!!!

There are 4 calibration spreadsheets in this ZIP file. The files correspond to:

- 5585 LS-1-MC Cal Cert.xls  - Luminance calibration (integrated over the spectrum). You can use this data to check linearity between your Photometer and the sphere.

- 5585 SR Full Radiance Cal Cert.xls – Spectral radiance calibration for the sphere with all the lamps on.

- 5585 SR ETHS-150 Cal Cert.xls -  Spectral radiance calibration for the sphere with all the only the ETHS-150 lamp on.

- 5585 Uniformity Cal Cert.xls  - Measurement of the spatial uniformity of the sphere

General instructions/recommendations:

  - The sphere should start with all the lamps on.

  - Lamps should be turned off one by one. Choose the best configuration in order to produce the best spread of points in your final plot.

  - Make sure the last lamp to be turned off is the ETHS-150

  - The luminance calibration should use the 4 points calibration

  - The spectral radiance calibration has only 2 points and should be interpolated for other light levels.

  - The spectral calibration should be integrated of the spectral response or your LED.