## Cover
From the Laboratories at
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel 603-715-3000 • Fax: 603-225-3089
CALIBRATION CERTIFICATE
This calibration is traceable to the National Institute of Standards and Technology in accordance with ISO 10012-1 and MIL-STD-45662. The calibration was accomplished by comparison to standards maintained by the laboratories at SphereOptics LLC. Complete compliance records and procedures are maintained by SphereOptics LLC and are available for inspection upon request.
CUSTOMER: | University of Maryland
1000 Hilltop Circle
BALTIMORE, MD 21250
USA
PURCHASE ORDER #: | UMBC1-0000005227
MODEL #: | SR-3BC
SERIAL #: | 5585
ITEM DESCRIPTION: | SPECTRAL RADIANCE CALIBRATION 300 - 2400nm
As Received |  | As Left
| New Equipment |  | In Tolerance*
| Out of Tolerance* |  | Out of Tolerance*
| Inoperative |  | Inoperative
| Limited Calibration** |  | Calibration
| Used Equipment |  | Adjusted
* Per Referenced calibration procedure
** Limited by customer of equipment performance
This report or any attached document within this report shall not be reproduced, except in full, without the written permission of SphereOptics LLC.
CALIBRATED: | 2008-07-29 | TECHNICIAN: | JJacobs
AUTHORIZED SIGNATURE:
RECALIBRATION DUE: | 2009-07-29
Calibration Lab Manager

## Data
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Fax: 603-225-3089
www.sphereoptics.com
CUSTOMER: |  | University of Maryland
MODEL NUMBER: |  | SR-3BC
SERIAL NUMBER: |  | 5585
DATE: |  | 2008-07-29
SPECTRAL RADIANCE DATA
EQUIPMENT USED
|  |  | Lamp | Address |  | Current (A) | Volts (V)
Agilent | 3634A |  | ETHS-150 | 5 |  | 5.79 | 19.52
Agilent | 5566 |  | ITHS-300 | 6 |  | 2.36 | 108.8
Agilent | 3634A |  | ITHS-150 | 7 |  | 6.37 | 24
Agilent | 3634A |  | ITHS-75 | 8 |  | 6.1 | 11.75
Kiethley | 6485 | [5 empty] | 0.0001655 | A
SOH-1 | Radiance Standard
* Above PS Address must be the same for LR software to work correctly
** Above readings taken at full luminance
Wavelength | Radiance |  |  | Wavelength | Radiance |  | Wavelength | Radiance
nm |  | uW/cm2-sr-nm |  | nm | uW/cm2-sr-nm |  | nm | uW/cm2-sr-nm
300 |  | 1.1755 |  | 580 | 78.3292 |  | 860 | 129.5709
310 |  | 1.6017 |  | 590 | 82.0227 |  | 870 | 129.2243
320 |  | 2.0662 |  | 600 | 85.701 |  | 880 | 128.5727
330 |  | 3.1624 |  | 610 | 89.1461 |  | 890 | 128.9369
340 |  | 3.8188 |  | 620 | 92.8547 |  | 900 | 128.9519
350 |  | 4.5735 |  | 630 | 96.2925 |  | 910 | 127.8862
360 |  | 6.1134 |  | 640 | 99.5095 |  | 920 | 127.397
370 |  | 7.6252 |  | 650 | 102.7034 |  | 930 | 126.6664
380 |  | 9.4603 |  | 660 | 105.4384 |  | 940 | 125.3284
390 |  | 11.5088 |  | 670 | 108.4052 |  | 950 | 124.7054
400 |  | 13.4239 |  | 680 | 110.9261 |  | 960 | 123.3607
410 |  | 16.3387 |  | 690 | 113.2043 |  | 970 | 123.0157
420 |  | 18.6132 |  | 700 | 115.8136 |  | 980 | 121.9623
430 |  | 21.2931 |  | 710 | 117.7046 |  | 990 | 121.1796
440 |  | 24.3597 |  | 720 | 119.7099 |  | 1000 | 120.546
450 |  | 27.573 |  | 730 | 121.3038 |  | 1010 | 119.6164
460 |  | 31.0838 |  | 740 | 122.6616 |  | 1020 | 118.5576
470 |  | 34.5794 |  | 750 | 124.6497 |  | 1030 | 117.738
480 |  | 38.3227 |  | 760 | 125.1878 |  | 1040 | 117.1142
490 |  | 42.128 |  | 770 | 126.1418 |  | 1050 | 115.8738
500 |  | 45.9879 |  | 780 | 127.2429 |  | 1060 | 114.5407
510 |  | 50.1553 |  | 790 | 127.9356 |  | 1070 | 113.1913
520 |  | 54.1367 |  | 800 | 129.1538 |  | 1080 | 111.7058
530 |  | 58.1757 |  | 810 | 129.0391 |  | 1090 | 110.4835
540 |  | 62.0428 |  | 820 | 129.4014 |  | 1100 | 109.151
550 |  | 66.0282 |  | 830 | 130.0296 |  | 1110 | 107.401
560 |  | 70.3014 |  | 840 | 129.8389 |  | 1120 | 105.4861
570 |  | 74.5651 |  | 850 | 129.6501 |  | 1130 | 103.8843
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Tel: 603-225-3089
Wavelength | Radiance |  |  | Wavelength | Radiance |  | Wavelength | Radiance
nm |  | uW/cm2-sr-nm |  | nm | uW/cm2-sr-nm |  | nm | uW/cm2-sr-nm
1140 |  | 102.0914 |  | 1590 | 40.7012 |  | 2040 | 11.5609
1150 |  | 100.1023 |  | 1600 | 40.2564 |  | 2050 | 11.2429
1160 |  | 98.3843 |  | 1610 | 39.7002 |  | 2060 | 11.2131
1170 |  | 96.9838 |  | 1620 | 39.2041 |  | 2070 | 11.0401
1180 |  | 95.2933 |  | 1630 | 38.6975 |  | 2080 | 10.9251
1190 |  | 93.623 |  | 1640 | 38.0689 |  | 2090 | 10.7591
1200 |  | 92.1878 |  | 1650 | 37.5533 |  | 2100 | 10.5659
1210 |  | 90.8401 |  | 1660 | 36.7947 |  | 2110 | 10.4199
1220 |  | 89.6387 |  | 1670 | 36.0671 |  | 2120 | 10.5954
1230 |  | 88.4676 |  | 1680 | 35.2787 |  | 2130 | 10.5771
1240 |  | 87.2759 |  | 1690 | 34.4495 |  | 2140 | 10.5986
1250 |  | 86.0265 |  | 1700 | 33.361 |  | 2150 | 10.3509
1260 |  | 84.8311 |  | 1710 | 32.1882 |  | 2160 | 10.2957
1270 |  | 83.5812 |  | 1720 | 31.4157 |  | 2170 | 10.0339
1280 |  | 82.2202 |  | 1730 | 30.7828 |  | 2180 | 9.7126
1290 |  | 80.5568 |  | 1740 | 30.0214 |  | 2190 | 9.4675
1300 |  | 78.9135 |  | 1750 | 29.0191 |  | 2200 | 9.4157
1310 |  | 77.3975 |  | 1760 | 28.1469 |  | 2210 | 9.1929
1320 |  | 75.8072 |  | 1770 | 27.8459 |  | 2220 | 9.0569
1330 |  | 73.7433 |  | 1780 | 27.3079 |  | 2230 | 8.8279
1340 |  | 71.2555 |  | 1790 | 26.5176 |  | 2240 | 8.8273
1350 |  | 68.8159 |  | 1800 | 25.875 |  | 2250 | 8.4477
1360 |  | 66.7116 |  | 1810 | 24.9818 |  | 2260 | 8.1095
1370 |  | 64.6117 |  | 1820 | 24.243 |  | 2270 | 7.7305
1380 |  | 61.2562 |  | 1830 | 23.7379 |  | 2280 | 7.5573
1390 |  | 56.3352 |  | 1840 | 23.124 |  | 2290 | 7.0088
1400 |  | 50.9573 |  | 1850 | 22.4921 |  | 2300 | 6.9122
1410 |  | 47.1752 |  | 1860 | 21.5116 |  | 2310 | 6.4965
1420 |  | 44.7865 |  | 1870 | 19.5447 |  | 2320 | 6.4833
1430 |  | 43.5235 |  | 1880 | 17.075 |  | 2330 | 6.3062
1440 |  | 42.5035 |  | 1890 | 14.0573 |  | 2340 | 6.0582
1450 |  | 41.7551 |  | 1900 | 11.4073 |  | 2350 | 5.8316
1460 |  | 41.5061 |  | 1910 | 10.0287 |  | 2360 | 5.5017
1470 |  | 41.5512 |  | 1920 | 9.4599 |  | 2370 | 5.264
1480 |  | 41.9826 |  | 1930 | 9.2166 |  | 2380 | 5.0095
1490 |  | 42.4848 |  | 1940 | 9.2522 |  | 2390 | 4.9578
1500 |  | 42.9821 |  | 1950 | 9.5943 |  | 2400 | 4.8608
1510 |  | 43.1842 |  | 1960 | 9.852
1520 |  | 43.2576 |  | 1970 | 10.2567
1530 |  | 43.2075 |  | 1980 | 10.6814
1540 |  | 43.033 |  | 1990 | 11.0758
1550 |  | 42.705 |  | 2000 | 11.4311
1560 |  | 42.35 |  | 2010 | 11.5423
1570 |  | 41.7694 |  | 2020 | 11.6681
1580 |  | 41.2122 |  | 2030 | 11.6669
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Tel: 603-225-3089
NIST CALIBRATION TRANSFER UNCERTAINTY
Transfer uncertainty of our secondary standards relative to the National Institute of Standards and Technology (NIST) scale is estimated as follows:
Wavelength (nm) |  | NIST Spectral Radiance Standard |  | SOH transfer uncertianty | Total Uncertianty
|  | k=2 (%) |  | k=2(%)
400 |  | 1.09 |  | 0.63 | 1.259
500 |  | 0.85 |  | 0.44 | 0.95713
800 |  | 0.68 |  | 0.33 | 0.75584
900 |  | 0.52 |  | 0.29 | 0.5954
1000 |  | 0.47 |  | 0.32 | 0.56859
1100 |  | 1.18 |  | 0.56 | 1.3061
1500 |  | 1.18 |  | 0.58 | 1.3148
2000 |  | 1.75 |  | 0.71 | 1.8885
2400 |  | 2.6 |  | 8.86 | 9.2336
This accuracy statement assumes similar geometric use of the standards by the customer and does not account for drift with time and use for the individual items issued.

## Raw
nm | radiance(uW/cm2-sr-nm)
300 | 1.1755
310 | 1.6017
320 | 2.0662
330 | 3.1624
340 | 3.8188
350 | 4.5735
360 | 6.1134
370 | 7.6252
380 | 9.4603
390 | 11.5088
400 | 13.4239
410 | 16.3387
420 | 18.6132
430 | 21.2931
440 | 24.3597
450 | 27.573
460 | 31.0838
470 | 34.5794
480 | 38.3227
490 | 42.128
500 | 45.9879
510 | 50.1553
520 | 54.1367
530 | 58.1757
540 | 62.0428
550 | 66.0282
560 | 70.3014
570 | 74.5651
580 | 78.3292
590 | 82.0227
600 | 85.701
610 | 89.1461
620 | 92.8547
630 | 96.2925
640 | 99.5095
650 | 102.7034
660 | 105.4384
670 | 108.4052
680 | 110.9261
690 | 113.2043
700 | 115.8136
710 | 117.7046
720 | 119.7099
730 | 121.3038
740 | 122.6616
750 | 124.6497
760 | 125.1878
770 | 126.1418
780 | 127.2429
790 | 127.9356
800 | 129.1538
810 | 129.0391
820 | 129.4014
830 | 130.0296
840 | 129.8389
850 | 129.6501
860 | 129.5709
870 | 129.2243
880 | 128.5727
890 | 128.9369
900 | 128.9519
910 | 127.8862
920 | 127.397
930 | 126.6664
940 | 125.3284
950 | 124.7054
960 | 123.3607
970 | 123.0157
980 | 121.9623
990 | 121.1796
1000 | 120.546
1010 | 119.6164
1020 | 118.5576
1030 | 117.738
1040 | 117.1142
1050 | 115.8738
1060 | 114.5407
1070 | 113.1913
1080 | 111.7058
1090 | 110.4835
1100 | 109.151
1110 | 107.401
1120 | 105.4861
1130 | 103.8843
1140 | 102.0914
1150 | 100.1023
1160 | 98.3843
1170 | 96.9838
1180 | 95.2933
1190 | 93.623
1200 | 92.1878
1210 | 90.8401
1220 | 89.6387
1230 | 88.4676
1240 | 87.2759
1250 | 86.0265
1260 | 84.8311
1270 | 83.5812
1280 | 82.2202
1290 | 80.5568
1300 | 78.9135
1310 | 77.3975
1320 | 75.8072
1330 | 73.7433
1340 | 71.2555
1350 | 68.8159
1360 | 66.7116
1370 | 64.6117
1380 | 61.2562
1390 | 56.3352
1400 | 50.9573
1410 | 47.1752
1420 | 44.7865
1430 | 43.5235
1440 | 42.5035
1450 | 41.7551
1460 | 41.5061
1470 | 41.5512
1480 | 41.9826
1490 | 42.4848
1500 | 42.9821
1510 | 43.1842
1520 | 43.2576
1530 | 43.2075
1540 | 43.033
1550 | 42.705
1560 | 42.35
1570 | 41.7694
1580 | 41.2122
1590 | 40.7012
1600 | 40.2564
1610 | 39.7002
1620 | 39.2041
1630 | 38.6975
1640 | 38.0689
1650 | 37.5533
1660 | 36.7947
1670 | 36.0671
1680 | 35.2787
1690 | 34.4495
1700 | 33.361
1710 | 32.1882
1720 | 31.4157
1730 | 30.7828
1740 | 30.0214
1750 | 29.0191
1760 | 28.1469
1770 | 27.8459
1780 | 27.3079
1790 | 26.5176
1800 | 25.875
1810 | 24.9818
1820 | 24.243
1830 | 23.7379
1840 | 23.124
1850 | 22.4921
1860 | 21.5116
1870 | 19.5447
1880 | 17.075
1890 | 14.0573
1900 | 11.4073
1910 | 10.0287
1920 | 9.4599
1930 | 9.2166
1940 | 9.2522
1950 | 9.5943
1960 | 9.852
1970 | 10.2567
1980 | 10.6814
1990 | 11.0758
2000 | 11.4311
2010 | 11.5423
2020 | 11.6681
2030 | 11.6669
2040 | 11.5609
2050 | 11.2429
2060 | 11.2131
2070 | 11.0401
2080 | 10.9251
2090 | 10.7591
2100 | 10.5659
2110 | 10.4199
2120 | 10.5954
2130 | 10.5771
2140 | 10.5986
2150 | 10.3509
2160 | 10.2957
2170 | 10.0339
2180 | 9.7126
2190 | 9.4675
2200 | 9.4157
2210 | 9.1929
2220 | 9.0569
2230 | 8.8279
2240 | 8.8273
2250 | 8.4477
2260 | 8.1095
2270 | 7.7305
2280 | 7.5573
2290 | 7.0088
2300 | 6.9122
2310 | 6.4965
2320 | 6.4833
2330 | 6.3062
2340 | 6.0582
2350 | 5.8316
2360 | 5.5017
2370 | 5.264
2380 | 5.0095
2390 | 4.9578
2400 | 4.8608