## Cover
From the Laboratories at
SphereOptics, LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel 603-715-3000 • Fax: 603-225-3089
CALIBRATION CERTIFICATE
This calibration is traceable to the National Institute of Standards and Technology in accordance with ISO 10012-1 and MIL-STD-45662. The calibration was accomplished by comparison to standards maintained by the laboratories at SphereOptics LLC. Complete compliance records and procedures are maintained by SphereOptics LLC and are available for inspection upon request.
CUSTOMER: | University of Maryland
1000 Hilltop Circle
BALTIMORE, MD 21250
USA
PURCHASE ORDER #: | UMBC1-0000005227
MODEL #: | LS-1-MC
SERIAL #: | 5585
ITEM DESCRIPTION: | Custom Calibration of Luminance
As Received |  | As Left
| New Equipment |  | In Tolerance*
| Out of Tolerance* |  | Out of Tolerance*
| Inoperative |  | Inoperative
| Limited Calibration** |  | Calibration
| Used Equipment |  | Adjusted
* Per Referenced calibration procedure
** Limited by customer of equipment performance
TEST RESULTS | See attached page
This report or any attached document within this report shall not be reproduced, except in full, without the written permission of SphereOptics, LLC.
CALIBRATED: | 2008-07-29 | TECHNICIAN: | JJacobs
AUTHORIZED SIGNATURE:
RECALIBRATION DUE: | 2009-07-29
Calibration Lab Manager

## Data
SphereOptics LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Tel: 603-225-3089
www.sphereoptics.com
CUSTOMER: |  | University of Maryland
MODEL NUMBER: |  | LS-1-MC
SERIAL NUMBER: |  | 5585
DATE: |  | 2008-07-29
EQUIPMENT USED
|  | Lamp | Address |  | Current (A) | Volts (V)
Agilent | 3634A | ETHS-150 | 5 |  | 5.79 | 19.52
Agilent | 5566 | ITHS-300 | 6 |  | 2.36 | 108.8
Agilent | 3634A | ITHS-150 | 7 |  | 6.37 | 24
Agilent | 3634A | ITHS-75 | 8 |  | 6.1 | 11.75
SOH | SMS-500 |  |  |  |  | 15256 | fL
SOH | SMS-500 |  |  |  |  | 3111 | K
Kiethley | 6485 |  |  |  |  | 0.00016698 | A
* Above PS Address must be the same for LR software to work correctly
** Above readings taken at full luminance
Lamp: | ITHS-300, 300W |  | Model : | Osram FKW 300W120V GY9.35
| ETHS-150, 150W |  |  | Osram HLX 64642 150W24V G6,35 XENOPHOT
| ITHS-150, 150W |  |  | Osram HLX 64642 150W24V G6,35 XENOPHOT
| ITHS-75, 75W |  |  | Gilway L7386A JC/GY6.35 75W12V2000H
Possible Luminance Levels
14120-15248fL | 600W: All-Lamps On Maximum (100%)
12182-13310fL | 525W: 300W, 150W and VA@100%
9326-10454fL | 450W: 300W, 75W, VA@100%
7422-8550fL | 375W: 300W, VA@100%
6567-7695fL | 300W: 150W, 75W, VA@100%
4681-5809fL | 225W: 150W, VA@100%
1893-3021fL | 150W: 75W, VA@100%
0-1128fL | 75W: VA@100%
Lamp Wattage Powered | Luminance (fL) | CCT (K) | Detector (A)
600 | 15248 | 3111 | 0.00016658
525 | 13310 | 3113 | 0.00014555
450 | 10454 | 3115 | 0.0001144
375 | 8550 | 3117 | 9.377e-05
300 | 7695 | 3114 | 8.437e-05
225 | 5809 | 3115 | 6.381e-05
150 | 3021 | 3108 | 3.296e-05
75 | 1128 | 3106 | 1.2334e-05
As VA Closes: | 1037 | 3105 | 1.1358e-05
| 868 | 3105 | 9.493e-06
| 686.9 | 3102 | 7.5234e-06
| 513.7 | 3097 | 5.6299e-06
| 344 | 3089 | 3.775e-06
| 192.1 | 3081 | 2.1057e-06
| 128.5 | 3075 | 1.4117e-06
| 77.05 | 3067 | 8.454e-07
| 40.04 | 3063 | 4.4e-07
| 17.54 | 3061 | 1.9263e-07
| 11.84 | 3061 | 1.3017e-07
SphereOptics, LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Tel: 603-225-3089
EQUATIONS USED IN LR SOFTWARE:
Luminance |  |  |  | Correlated Color Temperature
fL = A+BI+CI^2+DI^3 |  |  |  | CCT = E+FI+GI^2+HI^3
Group 1 (above 2.106e-6 Amps): |  |  |  | E = | 3087.676
A = | 0.53726 |  |  | F = | 2518950.6217
B = | 91305884.6493 |  |  | G = | -110008390945.585
C = | -3539505351.0356 |  |  | H = | 2277133171081090
D = | 30015441757184 |  |  | J = | -2.3219e+19
I = | current from detector |  |  | K = | 1.1336e+23
Group 2 (2.106e-6 to 1.302e-7 Amps): |  |  |  | L = | -2.1205e+26
A = | -0.091648 |  |  | I = | current from detector
B = | 91634685.8596
C = | -781847315930.3831
D = | 2.8909e+17
I = | current from detector
Group 3 (linear below 1.302e-7 Amps)
B = | 90957978.0287
I = | current from detector
SphereOptics, LLC
One Horseshoe Pond Ln
Concord, NH 03301
Tel: 603-715-3000
Tel: 603-225-3089
NIST CALIBRATION TRANSFER UNCERTAINTY
Transfer uncertainty of our secondary standards relative to the National Institute of Standards and Technology (NIST) scale is estimated as follows:
Luminous Intensity: |  |  |  |  | 1%
Color Temperature: |  |  |  |  | 0.5%
Luminous Flux (SCP): |  |  |  |  | 2%
Luminous Directional Transmittance: |  |  |  |  | 0.05%
Spectral Irradiance: |  |  | 350-1100 nm |  | 1%
Spectral Radiance: |  |  | 350-1100 nm |  | 1.5%
Spectral Transmittance: |  |  | 350-1200 nm |  | 0.1%
Detector Spectral Response: |  |  | 350-1100nm |  | 2%
Hoffman Engineering to SphereOptics:
|  |  | 360-460 nm |  | 10%
|  |  | 400-450 nm |  | 4%
|  |  | 450-800 nm |  | 1.5%
|  |  | 800-900 nm |  | 3%
|  |  | 900-1000 nm |  | 5%
This accuracy statement assumes similar geometric use of the standards by the customer and does not account for drift with time and use for the individual items issued.

## Raw
[5 empty] | Calculated | Calculated | Diff | Diff
| Watts | Detector (A) | Luminance (fL) | CCT (K) | Luminance (fL) | CCT (K) | Luminance (fL) | CCT (K)
| 600 | 0.00016658 | 15248 | 3111 | 15250.7979 | 3110.7019 | -0.01835% | 0.009581% | 3110.9457 | 0.001746% | y = 3.001544175718400E+13x3 - 3.539505351035600E+09x2 + 9.130588464929670E+07x + 5.372588110546060E-01
| 525 | 0.00014555 | 13310 | 3113 | 13307.6761 | 3113.4416 | 0.01746% | -0.01419% | 3113.2391 | -0.007681%
| 450 | 0.0001144 | 10454 | 3115 | 10444.5466 | 3115.5042 | 0.09043% | -0.01619% | 3114.319 | 0.02186%
| 375 | 9.377e-05 | 8550 | 3117 | 8555.9156 | 3115.5505 | -0.06919% | 0.0465% | 3116.5974 | 0.01292%
| 300 | 8.437e-05 | 7695 | 3114 | 7696.8459 | 3115.2209 | -0.02399% | -0.03921% | 3116.3153 | -0.07435%
| 225 | 6.381e-05 | 5809 | 3115 | 5820.1524 | 3113.7314 | -0.192% | 0.04073% | 3112.7831 | 0.07117%
| 150 | 3.296e-05 | 3021 | 3108 | 3007.2088 | 3109.507 | 0.4565% | -0.04849% | 3109.4627 | -0.04706%
| 75 | 1.2334e-05 | 1128 | 3106 | 1126.2219 | 3105.3425 | 0.1576% | 0.02117% | 3105.7764 | 0.0072%
|  | 1.1358e-05 | 1037 | 3105 | 1037.1769 | 3105.1187 | -0.01706% | -0.003823% | 3105.0658 | -0.002118%
|  | 9.493e-06 | 868 | 3105 | 867.0107 | 3104.6844 | 0.114% | 0.01017% | 3103.4428 | 0.05015%
|  | 7.5234e-06 | 686.9 | 3102 | 687.2804 | 3104.216 | -0.05538% | -0.07144% | 3101.2984 | 0.02262%
|  | 5.6299e-06 | 513.7 | 3097 | 514.4734 | 3103.7565 | -0.1506% | -0.2182% | 3098.7543 | -0.05664%
|  | 3.775e-06 | 344 | 3089 | 345.1681 | 3103.2975 | -0.3396% | -0.4629% | 3095.7352 | -0.218%
|  | 2.1057e-06 | 192.1 | 3081 | 192.0959 | 3102.8769 | 0.002133% | -0.7101% | 3092.5132 | -0.3737% | y = 2.890855657022420E+17x3 - 7.818473159303830E+11x2 + 9.163468585960700E+07x - 9.164779362820870E-02
|  | 1.4117e-06 | 128.5 | 3075 | 128.5242 | 3102.7 | -0.01883% | -0.9008% | 3091.0191 | -0.5209%
|  | 8.454e-07 | 77.05 | 3067 | 76.9922 | 3102.5547 | 0.07502% | -1.159% | 3089.7283 | -0.7411%
|  | 4.4e-07 | 40.04 | 3063 | 40.1009 | 3102.4502 | -0.152% | -1.288% | 3088.7632 | -0.8411%
|  | 1.9263e-07 | 17.54 | 3061 | 17.533 | 3102.3862 | 0.03993% | -1.352% | 3088.1572 | -0.8872%
|  | 1.3017e-07 | 11.84 | 3061 | 11.8238 | 3102.37 | 0.1366% | -1.352% | 3088.002 | -0.8821%
|  | 0 | 0 |  | 0 | [5 empty] | y = 9.095797802873170E+07x
CCT | y = -2.120486725444060E+26x6 + 1.133568479109540E+23x5 - 2.321907312624990E+19x4 + 2.277133171081090E+15x3 - 1.100083909455850E+11x2 + 2.518950621699880E+06x + 3.087675992312120E+03
Group 1 (above 2.106e-6 Amps):
A = | 0.53726
B = | 91305884.6493
C = | -3539505351.0356
D = | 30015441757184
I = | current from detector
Group 2 (2.106e-6 to 1.302e-7 Amps):
A = | -0.091648
B = | 91634685.8596
C = | -781847315930.3831
D = | 2.8909e+17
I = | current from detector
Group 3 (linear below 1.302e-7 Amps)
B = | 90957978.0287
I = | current from detector
Correlated Color Temperature
CCT = E+FI+GI^2+HI^3+JI^4+KI^5+LI^6
E = | 3087.676
F = | 2518950.6217
G = | -110008390945.585
H = | 2277133171081090
J = | -2.3219e+19
K = | 1.1336e+23
L = | -2.1205e+26
I = | current from detector